®

Optune® is a CE marked device.

For Investigational Use Only.

| Study | EF-41 (KEYNOTE D58) |
| --- | --- |
| Sponsor | Novocure GmbH Park 6 CH-6039 Root D4 Switzerland |
| Document number | PUM-SD-670 |
| Version: | 1 |

Table of Contents

1.	GLOSSARY OF MEDICAL TERMS	3

2.	INTENDED USE	4

3.	WHAT IS OPTUNE®?	4

4.	CONTRAINDICATIONS, WARNINGS AND PRECAUTIONS	5

5.	 ABOUT THE OPTUNE® TREATMENT KIT	8

6.	THE OPTUNE® DEVICE	9

7.	THE INE TRANSDUCER ARRAYS	10

8.	BEFORE YOU BEGIN	10

9.	REMOVING THE INE TRANSDUCER ARRAY FROM ITS PACKAGE	11

10.	PREPARING YOUR HEAD FOR INE TRANSDUCER ARRAY PLACEMENT	12

11.	PLACING THE TRANSDUCER ARRAYS ON YOUR HEAD	13

12.	CONNECTING THE INE TRANSDUCER ARRAYS TO THE DEVICE	15

13.	THE CONNECTION CABLE	16

14.	STARTING & STOPPING THE DEVICE	17

15.	CONNECTING & DISCONNECTING THE BATTERY	22

16.	CHARGING THE BATTERY	25

17.	USING THE PLUG-IN POWER SUPPLY	28

18.	DISCONNECTING FROM THE DEVICE	29

19.	CARRYING THE DEVICE	31

20.	GLOSSARY OF GRAPHIC SYMBOLS	32

21.	ENVIRONMENTAL CONDITIONS FOR OPERATION, STORAGE & TRANSPORTATION	35

22.	TROUBLESHOOTING	36

23.	ASSISTANCE & INFORMATION	39

24.	TRAVELING WITH OPTUNE®	40

25.	EMITTED RADIATION & ELECTROMAGNETIC COMPATIBILITY	41

# GLOSSARY OF MEDICAL TERMS

Cancer – abnormal cell division that spreads without control

Chemotherapy - drugs that should destroy or harm cancer cells

Clinical study – a research study that involves people

Contraindications - situations when a treatment should not be used

Device Support Specialist (DSS) – Novocure’s device technician who assists in supporting technical aspects of Optune®

Electric Field Generator (the device) - a portable device for delivering TTFields to the brain of patients

Glioblastoma (GBM) – a type of brain cancer

INE Transducer Array (also called transducer array) – adhesive bandages that hold insulated ceramic discs that deliver TTFields to the scalp.

Local – in one part of the body

Pembrolizumab – a drug used to treat cancer

Radiation – a treatment involving x-rays used to kill tumor cells

Steroids – when taken by mouth or IV (through the vein), a medication used to lower swelling around a brain tumor and help with symptoms related to the brain. When used on the skin, a medication that can reduce inflammation

Temozolomide (TMZ) – a type of cancer drug used to treat GBM

Topical – on the surface of the skin

TTFields – Tumor Treating Fields: Alternating electric fields, delivered using transducer arrays to the part of the body with a solid tumor. The fields have been shown to destroy tumor cells.

Tumor – an abnormal growth of tissue

# INTENDED USE

In this clinical study Optune® concomitant with temozolomide and pembrolizumab/placebo is intended for the treatment of newly diagnosed, supratentorial glioblastoma following maximal debulking surgery and completion of radiation therapy (RT) together with concomitant standard of care chemotherapy.

# WHAT IS OPTUNE®?

Optune® is a portable, battery operated device which delivers electric fields, called Tumor Treating Fields (TTFields) to the patient by means of insulated INE Transducer Arrays.

TTFields produces electric forces intended to destroy brain cancer cells. The device and battery are carried in a shoulder or a side bag.

Optune® is intended for treatment at least 18 hours a day on a monthly average.

# CONTRAINDICATIONS, WARNINGS AND PRECAUTIONS

Contraindications

Do not use Optune® if you are under 18 years old.

Do not use Optune® if you have an implantable electronic medical device, a skull defect (such as, missing bone with no replacement) or bullet fragments. Examples of active electronic devices include deep brain stimulators, spinal cord stimulators, vagus nerve stimulators, pacemakers, defibrillators, and programmable shunts. Use of Optune® together with implanted electronic devices has not been tested and may lead to malfunctioning of the implanted device. Use of Optune® together with skull defects or bullet fragments has not been tested and may possibly lead to tissue damage or render Optune® ineffective.

Do not use the Optune® treatment kit if you are known to be sensitive to conductive hydrogels like the gel used on electrocardiogram (ECG) stickers or transcutaneous electrical nerve stimulation (TENS) electrodes. In this case, skin contact with the gel used with the Optune® treatment kit may commonly cause increased redness and itching, and rarely may even lead to severe allergic reactions such as shock and respiratory failure.

Do not use the Optune® treatment kit if you are pregnant or breast feeding.

Warnings

Warning - Use Optune® only after receiving training from qualified personnel.

Warning - In case of skin irritation, which appears as redness under the INE Transducer Arrays (a mild rash), use high potency topical steroids when replacing transducer arrays. This will help relieve your skin irritation. If you do not use this cream, the skin irritation can become more serious and may even lead to skin break down, infections, pain and blisters. If this happens, stop using the topical steroid cream and contact your doctor. Your doctor will supply you with an antibiotic cream to use when replacing the transducer arrays. If you do not use this cream, your symptoms may continue, and your doctor may ask you to take a break from treatment until your skin heals. Taking a break from treatment may lower your chance to respond to treatment.

Warning - All servicing procedures must be performed by qualified and trained personnel. If you attempt to open the device alone, you may cause damage to the device. You could also get an electric shock by touching the inner parts of the device.

Precautions

Caution – Keep Optune® out of the reach of children. If children touch the device, they could damage the device. This could cause a break in treatment. Breaks in treatment may lower your chance to respond to treatment.

Caution - Do not use any parts that do not come with the Optune® treatment kit, or that were not sent to you by the device manufacturer or given to you by your doctor. Use of other parts, manufactured by other companies or for use with other devices, can damage the device. This may lead to a break in treatment. Breaks in treatment may lower your chance to respond to treatment.

Caution – If your doctor used plates or screws to close your skull bone during your surgery, be careful when placing the transducer arrays.  Make sure the round disks that make up the transducer arrays are not on top of the areas where you can feel the screws or plates under your skin.  In other words, make sure the screws or plates under your skin are in between the round disks that make up the transducer arrays.  If you do not do this, you may have increased skin damage which may lead to a break in treatment.  Breaks in treatment may lower your change to respond to treatment.

Caution – Tell your doctor before using the device if you have an inactive implanted medical device in your brain (such as a stent, plastic drug delivery reservoir, aneurysm clip or coil or device lead).  Use of Optune® in subjects with inactive implanted medical devices in their brain was not tested and could lead to tissue damage or lower your change to respond to treatment.

Caution – Do not use Optune® if any parts look damaged (torn wires, loose connectors, loose sockets, cracks or breaks in the plastic case). Use of damaged components can damage the device and cause a break in treatment. Breaks from treatment may lower your chance to respond to treatment.

Caution – Do not wet the device or transducer arrays. Getting the device wet may damage it, preventing you from receiving treatment for the right amount of time. Getting the transducer arrays very wet is likely to cause the transducer arrays to come loose from your head. If this happens, the device will turn off and you will need to change the transducer arrays.

Caution – Before connecting or disconnecting the transducer arrays, make sure that the Optune® power switch is in the OFF position. Disconnecting transducer arrays with the device power switch in the ON position may cause a device alarm to go off and could damage the device.

Caution - If you have an underlying serious skin condition on the scalp, discuss with your doctor whether this may prevent or temporarily interfere with Optune® treatment.

Notices

Notice! The Optune® device and transducer arrays will activate metal detectors.

Notice! You should use Optune® for at least 18 hours a day to get the best response to treatment. Using Optune® for less than 18 hours a day lowers the chances that you will respond to treatment.

Notice! Do not stop using Optune® even if you have used it less than the recommended 18 hours per day. You should stop using the device only if your doctor tells you to. Stopping treatment could lower the chances that you will respond to treatment.

Notice! If you plan to be away from home for more than 2 hours, carry an extra battery and/or the power supply with you in case the battery you are using runs out. If you do not take a spare battery and/or the power supply, you may have a break in your treatment. Breaks in treatment may lower your chance to respond to treatment.

Notice! Make sure you have at least 12 extra transducer arrays at all times. This will last you until the next transducer array shipment arrives. Remember to order more transducer arrays when there are at least 12 extra transducer arrays left. If you do not order transducer arrays in time you may have a break in your treatment. Breaks in treatment may lower your chance to respond to treatment.

Notice! Batteries may weaken over time and need to be replaced. You will know this has happened when the amount of time the device can run on a fully charged battery begins to shorten. For example, if the low battery indicator light flashes within only 1.5 hours from the start of treatment, replace the battery. If you do not have replacement batteries when your batteries run out, you may have a break in your treatment. Breaks in treatment may lower your chance to respond to treatment.

Notice! You should carry the Troubleshooting Guide (Section ‎22) at all times.  This guide is necessary to ensure Optune® works properly.  If you do not work the device correctly, you may have a break in your treatment.  Breaks in treatment may lower your chance to respond to treatment.

Notice! Do not block the device vents located on the front and back of the Electric Field Generator. Blocking the vents may cause the device to overheat and turn off, leading to a break in treatment. If this happens, unblock the vents, wait 5 minutes and restart the device.

Notice! Do not block the battery charger vents located on the sides of the battery chargers. Blocking the vents may cause the charger to overheat. This could prevent your batteries from charging.

Notice! Before using a transducer array make sure its package is sealed by gently rubbing the package between thumb and pointer finger on all four sides. The package should be closed on all sides. There should be no openings in the package seal. If the package is not sealed, the transducer array may be damaged. A damaged transducer array will not work properly and may cause the device to turn off.

Notice! The transducer arrays are for single use and should not be taken off your head and put back on again. If you put a used transducer array back on your head again, it may not stick well to your skin and the device could turn off.

# ABOUT THE OPTUNE® TREATMENT KIT

| Optune® treatment kit | Model Number |
| --- | --- |
| 1 Electric Field Generator (the device) | TFH9100 |
| 2 Battery | IBH9100 |
| 3 Charger for Battery | ICH9100 |
| 4 Plug In Power Supply | SPS9100 |
| 5 Connection Cable and Box (CAD) | CAD9100 |
| 6 INE Transducer Arrays | INE9TAN and INE9TANW |
| 7 Power Cords | NA |
| 8 Shoulder Bag and Strap | NA |
| 9 Battery Case | NA |

# THE OPTUNE® DEVICE

The Optune® is a portable medical device. It delivers electric fields called “TTFields” to the brain using INE Transducer Arrays.  Keep the TTFields treatment on all the time, as much as possible, when awake and when sleeping.  Keep breaks from treatment as short as possible.  The picture below shows the device controls to work the system.

You do not need to adjust any settings. You only need to turn the device and the therapy off and on.

The device has the following controls:

- Power Switch

- Power Supply Port

- TTFields Therapy ON/OFF Button

- Power ON/ Battery/ Error Indicators

- Connection Cable (CAD) Port

- Battery Gauge

- Battery Charge Indicator Button

# THE INE TRANSDUCER ARRAYS

The INE Transducer Arrays are arrays of insulated transducers which deliver Tumor Treating Fields (TTFields) to a tumor non-invasively.

The INE Transducer Arrays are supplied sterile and are to be used with the Optune® treatment kit only.  The INE Transducer Arrays do not require maintenance, cleaning or disinfection.

# BEFORE YOU BEGIN

You will need to use four (4) transducer arrays (two black and two white arrays) each time you change the arrays. The black arrays are placed on the front and back of your head, and the white arrays are placed on the sides of your head. Remember: Black goes on the back, white goes on the right. Change the four (4) transducer arrays at least two times per week (every 4 days at most) to continue treatment with Optune®.

You may change the transducer arrays with the help of a doctor, a nurse or caregiver if needed.

# REMOVING THE INE TRANSDUCER ARRAY FROM ITS PACKAGE

Open the see-through envelope of four (4) INE Transducer Arrays by gently pulling apart the opposing edges of the envelope. Hold the INE Transducer Array as shown in the picture.

# PREPARING YOUR HEAD FOR INE TRANSDUCER ARRAY PLACEMENT

Wash your hands thoroughly and wash your head with a gentle shampoo.

If this is the first time you have used the transducer arrays, ignore this step and skip ahead to the next step (shaving).  If you are replacing transducer arrays, you, or your doctor or caregiver if needed, should wipe the skin with baby oil to remove any old adhesive from other transducer arrays. It will not stop the device from working.

Shave your entire scalp using an electric shaver. Do not leave any stubble.

Wipe your scalp with 70% Alcohol.  If the skin is irritated, avoid applying alcohol and wash your head with a gentle shampoo and water before gently patting it dry, using a soft towel.

If your scalp is red, apply the steroid cream prescribed to you by your physician.

If you have any sores on your scalp treat them as instructed by your treating physician.

Wait at least 15 minutes and gently clean your skin again with 70% alcohol to remove any residues.

Apply the transducer arrays after your scalp is dry.

# PLACING THE TRANSDUCER ARRAYS ON YOUR HEAD

After you prepare your scalp, put the INE Transducer Arrays on your head with the help of a doctor or a caregiver if needed. At least two times per week (every 4 days at most) remove the INE Transducer Arrays, prepare the scalp and put on a new set of the transducer arrays.

You will know it is time to change transducer arrays when the device alarm beeps more often. This means that the device is not able to work properly because of hair growth. Hair growth keeps the transducer arrays from making good contact with your scalp.

To place the INE Transducer Arrays on your head, with the help of a caregiver or doctor if needed, follow the steps below.

Note -if this is the first time you have used the INE Transducer Arrays, ignore the first step (removal).

  - Remove the transducer arrays from your head by peeling the medical tape away from your scalp.

  - Notice which color transducer array goes where on your head. The two black arrays are placed in the front and back of the head and the two white arrays are placed on the sides.

  - Prepare your skin for the transducer arrays, as described above.

  - Peel off the white layer (liner) covering the gel from the first transducer array.

If this is the first time you have used the transducer arrays, put the transducer arrays on your head as shown in the Transducer array layout (“map”) that your doctor gave you.

Placement is based on the location of your tumor. When changing the transducer arrays, place the transducer arrays on your head in the same general location as before, but shift the transducer arrays less than an inch in the direction of the arrow on your Transducer Array Layout or “map.” To reduce skin irritation under the transducer arrays, move the transducer arrays a small amount. Shifting the transducer arrays is not required for the device to work properly.

The ceramic gel pads leave a light impression on the surface of the scalp. This will help to position the new transducer arrays. The transducer arrays work in pairs, so they should always be moved together (see figure). The next time the transducer arrays are replaced, move them back to the previous position.

Patient 1	Patient 2	Patient 3

Place the other three transducer arrays in the same fashion.

Pull the tabs on each side of the transducer arrays and press them firmly to your scalp.

You may need to ask for assistance from a friend or family member to place the back transducer array.

Press the entire rim of the INE Transducer Array tape to your skin.

# CONNECTING THE INE TRANSDUCER ARRAYS TO THE DEVICE

Connect the four black and white INE Transducer Array connectors to the corresponding black and white coded sockets on the Optune® connection cable. Plug the transducer array connectors with the black connection ends into the two black sockets (there will be one labeled “P1” and one labeled “N1”) and the two white connection ends into the white sockets (there will be one labeled “P2” and one labeled “N2”) (see diagram).

Press firmly to verify the connectors are inserted all the way.

Collect the INE Transducer Array wires together and bind them with a small piece of tape where convenient. You may clip the connection cable clip to your belt.

# THE CONNECTION CABLE

The connection cable is the coiled, stretchy cord that runs from the device to the connection box. The four black (2) and white (2) transducer array connectors plug into the connection box. The black and white coding matches with the transducer array position on the head.

Plug the connection cable into the device in the socket on the left of the front panel. The connection cable plugs into the socket with the arrows facing up.

Push in the connector until you hear a click. The click means it is secure.

# STARTING & STOPPING THE DEVICE

In order to Start the Device and Begin Treatment:

Put the transducer arrays on the scalp (with the help of a caregiver if needed) as described above in section ‎0.

Plug the transducer arrays into the connection cable box, as described above in section ‎12.

Plug the connection cable into the device with the arrows on the connector up (as described in section ‎13).

Insert a charged battery into the device (see Section ‎15) or attach the wall power supply (Section‎17).

Turn the power button to the ON position.  The power button is on the bottom of the device.

Allow about 10 seconds for the self-check to be completed. The “Power” indicator on the front panel of the device will illuminate green.

If a charged battery is installed and there is no wall power supply plugged in, the "Battery" indicator will also illuminate green.

If a wall power supply is plugged into the device, the device will run using the wall power supply and the "Battery" indicator will not illuminate.

Press the TTFields therapy button once – this will start treatment.

The “TTFIELDS” above the TTFields therapy button will illuminate blue and stay on while the treatment is on. If the blue words are not illuminated, then the treatment is not running, and you should check the setup and restart the procedure.

If, after this, the indicator lights do not light up, consult the Troubleshooting Guide (Section ‎22). If you still have problems, contact the technical support or your DSS (Section ‎23).

Note: The green, blue and yellow indicators will dim in a dark room and will brighten in a light environment. The red error indicator light will not be dimmed in any case.

If the therapy button is not pressed within 10 minutes after the device is turned ON, an alarm will sound, indicating that the device is ON but the therapy is OFF. This is a reminder to start the therapy. The therapy button should be pressed once to silence the alarm and again to start the therapy. “TTFIELDS” will be illuminated blue when the patient is receiving treatment.

Turning Off the Device and Stopping Treatment

You may stop treatment if the following happens:

- If the Device is Running Properly, But You Need to Stop Treatment to Take a Break:

Press the TTFields button. The blue “TTFIELDS” lights around the button will turn off. This turns the TTFields therapy off, but the device power is still on.

Then, turn off the device by turning the power button on the bottom of the device to the OFF position.

- If an Error occurs:

If an error occurs, the device will turn off the TTFields and make a loud beeping noise.

The red Error light will light up (as shown below). In this case, once the red Error light illuminates, the treatment has stopped.

To turn off the device:

- Press the TTFields button on the front of the device to stop the alarm. The red Error light will turn off. (If this does not silence the alarm, proceed to the next step of turning off the power switch to silence the alarm. A small number of alarm conditions cannot be silenced by pushing the TTFields button).

- Turn off the device by turning the power button to the OFF position.

- See the Troubleshooting Guide (Section ‎22) for instructions on fixing problems.

Restart the device and restart treatment if no problem is found. If the alarm continues, contact your technical support or DSS (Section ‎23).

- If the Low Battery Indicator Light Lights Up

When your battery runs out (after about 2-3 hours), an alarm will beep, the TTFields therapy will stop and both the yellow low Battery light and red Error light will light up. This alarm sound is the same alarm sound the device makes for an error. However, in this case both the yellow and red lights will light up instead of just the red light.

To turn off the device:

- Press the TTFields button on the front of the device to stop the alarm. The red Error and the yellow low Battery lights will turn off.

- Replace the battery using the steps in Section ‎15.

# CONNECTING & DISCONNECTING THE BATTERY

The Optune® treatment kit comes with 4 rechargeable batteries. Batteries slide into the device. The battery should be inserted until you hear a “click”, indicating the battery is in place. Take care not to drop the battery in place or force it into the battery slot. The Optune® uses one battery at a time. The other three batteries should stay in the battery charger. Each battery lasts 2 to 3 hours. Replace the battery each time it runs out (when the yellow low Battery indicator is on, as described in section ‎14). If you plan to be away from home for more than 2 hours, carry extra batteries or a power supply.

Gently press down to lock the battery in place. Make sure the battery latch is fully engaged

To remove the battery from the slot, press both blue buttons on the side of the battery and lift up.

Recharge the batteries in the charger (see Section ‎8) for two to four hours. The batteries will keep most of their charge after being removed from the charger for several days but eventually will lose their charge.

It will not hurt the batteries to keep them in the charger after they are fully charged so you can leave them there if they are not needed.

You can charge and use the batteries many times for about six to nine months.

Over time, the length of time that the batteries can run the device (before the yellow low battery indicator light illuminates and the alarm beeps) will get shorter.

If the time from treatment start with a full battery to low battery alarm falls below 90 minutes, contact your technical support or DSS (see Section ‎23) to get replacement batteries.

The battery light will turn from green to yellow when the battery charge falls below a threshold. This is an indication that the battery should be changed soon.

The treatment will continue to run while the yellow low battery indicator is illuminated until the audible alarm sounds and the red error light illuminates.

Once this happens the treatment will stop, and the device must be turned off and the battery replaced.

When the yellow low Battery indicator light lights up, there are two ways you can replace the depleted battery with a charged battery.

Option One - (to be used if near the direct wall power supply) allows you to change the battery without interrupting therapy. This can be used before the battery is completely depleted, and before the device has alarmed.

Please follow these steps:

- Plug in the wall power supply to the bottom of the Optune® device. (See Section ‎17). The lights on the display panel will indicate you are no longer running on battery power.

- Remove the battery from the battery slot by pressing the blue buttons on the side of the battery and lifting the battery cartridge out of the slot.

- Select a fully charged battery from the charger or device bag.

- Slide the fully charged battery in the battery slot, gently push down to lock the battery in place.

- Remove the power cord from the bottom of the device.

The battery light on the display panel will illuminate indicating you are no longer running on wall power.

See the next section to check the battery gauge.

Option Two - If you are not near the wall power supply, or if the battery has totally depleted please replace the battery by following these steps:

- Press the TTFields button to stop the treatment.

- Turn off the device using the power switch (on the bottom of the device).

- Remove the battery from the battery slot by pressing the blue buttons on the side of the battery and lifting the battery cartridge out of the slot.

- Select a fully charged battery from the charger.

- Slide the fully charged battery in the battery slot, gently push down to lock the battery in place.

- See the next section to check the battery gauge.

- Turn ON the device using the power switch and wait for the device to run a self-check (this takes about 10 seconds). To start treatment, press the TTFields button.

- Insert the used battery into the charger for recharging (see Section ‎8).

# CHARGING THE BATTERY

Checking the Battery Gauge

While you are using the Optune® treatment kit, you may want to check how much energy is left in your battery. Checking the battery will not interfere with or stop your treatment. To check the battery life, press the button once on the top of the battery cartridge. The battery life will be indicated by the lighted gauge to the right of the button. The gauge reads from full to empty, like a gas gauge in a car.

Full charge

75% charge

50% charge

25% charge

Empty

Charging Batteries

The battery charger recharges used batteries. The battery charger uses power from a standard wall outlet. Each battery sits in a slot that connects it directly to the charger.

Before charging the batteries, plug the charger power cord into a standard wall outlet and turn on the power button at the back of the charger.

The front lights of the charger will come on during a self-check then the small light in the center of the front panel will light up green indicating power is applied.

To recharge a used battery:

- Place the used battery in one of the three openings in the top of the charger. Slide the battery in until it is fully in place.

- The light directly in front of the opening where the battery is plugged in will illuminate flashing green. This indicates the battery is charging. The green light will flash faster once the battery has been charged to 95% of its capacity. You can also check the battery gauge while charging to get information regarding the amount of charge in the battery.

- When the battery is fully charged (about 2 to 4 hours), the charge light will turn from flashing green to solid green. The solid green light will disappear upon removal of the battery or the disconnection of the charger from the standard wall outlet.

- Charger Power Cord

- Power Switch

Back View of the Battery Charger Showing the Power Switch

and Where the Power Cord Connects

If a light on the front panel turns red, this indicates that there is a fault with the battery or charger, and you should contact technical support or your DSS for assistance. Do not use a battery if it creates a red light on the charger.

Keep the batteries in the charger even after they are fully charged. This will not harm the batteries.

- Battery Charging Slot

- Battery Charge Indicator

- Charger Power Indicator

Front view of the battery charger showing how the batteries are inserted into the charger

Notice: The charger is not intended for use in the presence of flammable mixtures

# USING THE PLUG-IN POWER SUPPLY

When you plan to stay in one place for a while, like when you are sleeping, you may use the plug-in power supply instead of the batteries. Unlike the batteries, there is no limit to how long the device can work when you use the plug-in power supply. The plug-in power supply will work with either U.S. (120V AC) or European (230V AC) outlets.

Note: It is normal for the power supply to become warm when in use. If the power supply becomes too hot to touch, unplug it and contact technical support or your DSS (Section ‎23).

When the device has a battery and is also connected to the wall power supply, the device will utilize the wall power supply as the preferred power source. When the wall power cord is plugged in while the device is running on battery, the device will automatically switch from battery power to wall supply power.

### Connecting the Plug-In Power Supply

- Plug in the power supply to a standard wall outlet using the power cord that comes with it.

- You do not need to remove the battery from the device to use the wall power supply. Please note that a battery in the device will not charge while the device is plugged into the wall power supply. Depleted batteries must be placed on the battery charger to recharge. If the TTFields are activated you do not need to turn them off to plug in the wall power supply.

- Plug the connector of the wall power supply into the AC port on the back panel of the device (next to the power switch).

- If the TTFields are running, the device will automatically switch to wall power supply without interruption of the TTFields. If the device is not turned on, turn on the power switch and wait for the self-check to be completed (about 10 seconds). Push the TTFields button to start the device (as described in Section ‎14).

### To Disconnect the Plug-In Power Supply and Go Back to Battery Power

- Ensure that a charged battery is properly inserted in the device before removing the wall power supply. If the TTFields are running, you do not need to turn them off before removing the wall power supply. The device will automatically switch to battery power once the power supply is removed.

- Remove the connector of the plug-in power supply from the bottom panel of the device.

- If the device is not turned on, turn on the power switch and wait for the self-check to be completed (about 10 seconds). Push the TTFields button to start the device.

- Store the plug-in power supply for future use.

# DISCONNECTING FROM THE DEVICE

There are two ways to unplug from the device to take a break from treatment (after turning off the device)

- Unplug the connection cable from the device.

- Unplug the transducer arrays from the connection cable.

To Unplug the Connection Cable from the Device

- Stop TTFields treatment by pressing the TTFields button.

- Turn off the Optune® device using the power switch.

- Unplug the connection cable from the socket by holding the gray collar and pulling back.

CAUTION! Do not pull on the cord!

You may now move around without the device, but you will still be connected to the connection cable and box.

To start treatment again after your break:

- Plug the connection cable into the device socket with the arrows pointing up.

- Turn on the device using the power switch. Wait for self-check to be completed (about 10 seconds).

- Turn on the TTFields using the TTFields button

To Unplug the transducer arrays from the connection cable box

To take a break from treatment and completely disconnect from the device, unplug the transducer arrays from the connection cable box.  The connection cable remains plugged in the device socket.

- Stop treatment by pressing the TTFields button.

- Turn off the Optune® device using the power switch.

- Unplug the transducer array connectors from the connection box by pulling the connectors.

Note: You may have to wiggle the transducer array cables to remove them.

To restart treatment, plug the transducer arrays into the connection box. Plug each transducer array into its matching color (black or white) that goes with the transducer array’s position on the head (as described in Section ‎0).

- When all 4 transducer arrays are plugged in, turn on the device power switch and wait for self-check to be completed (about 10 seconds). Push the TTFields button to restart treatment.

# CARRYING THE DEVICE

Both the electric field generator and the battery fit in a carrying bag. The bag can be carried in one of two ways: by the handle on top or over the shoulder/cross-body with a carrying strap attached.

Note: Do not place the device in a different bag. The Optune® device has a fan on the inside that needs air flow. The bag that comes with the device is designed to allow for proper air flow. If you put the device in a bag without proper air flow, it could overheat and stop the treatment. If this happens, you will hear an alarm.

# GLOSSARY OF GRAPHIC SYMBOLS

|  | Caution Consult the instructions for use for important cautionary information such as warnings and precautions. |
| --- | --- |
|  | Date of Manufacturing |
|  | Fragile – handle with care |
|  | Follow instructions for use |
|  | Keep dry. Do not enter rooms with high humidity or danger of direct exposure to water while wearing the device. Do not use the device if not within its carrying bag. Do not expose the device to direct rain. |
|  | For indoor use only The charger and power supply are for indoor use only. |
|  | Waste Electrical and Electronic Equipment recycling “WEEE disposal” Contact technical support to arrange for proper disposal of arrays that are used up or no longer in use. |
|  | Batteries are Lithium Ion. Contact technical support to arrange for proper disposal of batteries that are used up or no longer in use. |
|  | Protect from heat and radioactive sources Optune® should be kept away from extreme heat and sources of radiation. |
|  | BF type applied part – symbolizes the part which comes in contact with the patient Specifies the P/N of the applied part to be used with this device. |
|  | Expiration date – do not use beyond this date |
|  | Power ON / OFF switch for the electric field generator and battery charger: When the switch is in the I position the device is ON. When the switch is in the O position the device is OFF. |
|  | Do not use if package is damaged Do not use the Transducer arrays if their packaging is breached. |
|  | Do not re-use The INE Transducer Arrays are for single use and should not be re-used. |
|  | Do not re-sterilize |
|  | Sterile/sterilization method The INE Transducer Arrays are sterilized by Gamma irradiation. |
|  | Model number |
|  | Catalogue number |
|  | Serial number |
|  | Lot number |
|  | Class II equipment per IEC 60601-1 |
| IP21 | Protects persons against access to hazardous parts with fingers. Protects the equipment inside the enclosure against ingress of solid foreign objects of 12.5 mm in diameter or greater. Protects the equipment inside the enclosure against ingress of vertical falling water drops. |
| IP22 | Protects persons against access to hazardous parts with fingers. Protects the equipment inside the enclosure against ingress of solid foreign objects of 12.5 mm in diameter or greater. Protects the equipment inside the enclosure against ingress of vertical falling water drops when enclosure is tilted up to 15°. |
|  | Storage Temperature range Do not expose to temperatures below -5°C or above 40°C. |
|  | Storage humidity range Do not expose to humidity below 15% or above 93%. |
|  | P1 P2 N1 N2 black and white coding on the connection box. |
|  | Manufacturer information: Novocure GmbH, Business Village D4, Park 6/Platz 10, 6039 Root, Switzerland. |

# ENVIRONMENTAL CONDITIONS FOR OPERATION, STORAGE & TRANSPORTATION

Conditions for operation

All treatment kit components shall be normally used under conditions specified below:

- Mainly for home use

- For indoor use only (chargers and power supply)

- Not for use in shower, bath tub or sink, or in heavy rain

- Not for use in presence of flammable mixtures

- Can be dropped on the floor with no resulting safety hazard, but is no longer expected to function

Conditions of visibility:  any

Cleaning

All treatment kit components can be periodically cleaned with a damp cloth to remove dust and regular dirt.

Physical operation conditions for all treatment kit components:

- Temperature range: -5°C/+40°C (23°F/104°F)

- Relative Humidity range: 15-93%

- Ambient pressure range: 700-1060hPa

Conditions for storage

- Temperature range: -5°C/+40°C (23°F/104°F) for the device and additional parts

- Temperature range: +5°C/+27°C (41°F/80°F) for the INE Transducer Arrays

Conditions for transport

Transportation of the device, additional parts and INE Transducer Arrays shall be possible using air/ground transportation in weather protected conditions as specified below:

- Temperature range: -5°C/+40°C (23°F/104°F)

- Maximal relative humidity 15-93%

- No direct exposure to water.

# TROUBLESHOOTING

| Problem | Possible Causes | Actions to be Taken |
| --- | --- | --- |
| Device power indicator light does not light up after turning ON the device | - Device not connected to power source - If Battery – battery depleted - Battery malfunction - If power supply – not properly plugged into the wall - Device malfunction - Power supply malfunction | - If on battery – check battery gauge to verify it is not depleted. If it is – replace with a charged battery or connect to the power supply - Verify both the device and the power source are properly connected and re-try - Evaluate the integrity of all connectors. Nothing should appear to be damaged or broken in any way - If device cannot be powered on by either the battery or the wall power supply or if anything appears to be damaged do not use the device - Contact your DSS for technical support |
| Any cable detached from transducer array/ connection cable/ device | - Too much physical force to cables - Device malfunction | - Silence the notification signal by pressing the TTFields button - Evaluate the connectors. If intact – reconnect and re-start therapy - If anything appears damaged or cannot be properly connected do not try to use the device - Contact your DSS for technical support |
| Device dropped or wet | Incorrect use | - Press TTFields button to stop therapy - Turn OFF power switch - Contact your DSS for technical support |
| Device alarm on | - Low battery - Cable loose or disconnected - Device is too hot – vents on the device are blocked - Local hot spot on transducer array from laying on a pillow, for example - Poor transducer array contact due to hair growth or other reason - Device malfunction - Device is turned ON, but the therapy has not been activated | If low Battery light is yellow: - Replace battery as described above in Section ‎8 - Turn on treatment If the Error light lights up but the low Battery light is not lit: - Press the TTFields button to stop the alarm - Wait a few seconds then press the TTFields button again - If the blue lights around the TTFields button light up – the therapy has now been activated If the notification signal recurs within a few minutes: - Silence the notification signal and power the device down completely - Disconnect all equipment and make sure that nothing appears to be damaged or broken. If something is – replace the damaged item before trying to power the device back - Re-connect all equipment in proper order and power the device back up. Verify the self-test is completed and press the TTFields button - Check vents on device to make sure they are not blocked - If lying down, move your head - Make sure transducer arrays are well stuck to the head, add tape if needed - Restart treatment - If alarm keeps going, turn off the device and contact your DSS |
| Notification alarm sounds 10 minutes after the device was powered on | 1. Therapy Timeout | The notification alarm on the device will sound if it is powered on for 10 minutes, but therapy is not initiated. This is a reminder to start therapy and does not indicate a malfunction. - Silence the notification alarm by pressing the TTFields button then wait a few seconds and press the TTFields button again to initiate treatment. The blue indicator around the TTFields button will illuminate to indicate therapy is now on - If you encounter further alarms please review the general "notification signal" section below |
| Low Battery indicator light remains on after battery replaced | - Charger malfunction - Battery malfunction - Device malfunction | - Replace battery with an additional charged battery - If problem is not fixed – contact your DSS for technical support |
| When powering on the device a continuous notification alarm sounds and all lights remain on indefinitely. Device does not complete the self-test | - Device is too hot - Device malfunction - Power Source Malfunction | - Power the device off completely using the power switch - Verify the device is not hot to the touch - Connect the device to a different power source and try powering on again - If device cannot be powered on by either the battery or the wall power supply or if anything appears to be damaged, please contact your DSS for technical support |
| Redness of the skin under the transducer arrays | Common side effect | Use over the counter 0.1% hydrocortisone cream when switching transducer arrays. Shift transducer arrays 3/4 of an inch from the last location (so the adhesive gel is between the red marks). If the redness gets worse: - See your study doctor |
| Blisters under the transducer arrays | Rare side effect | See your study doctor for a prescription antibacterial cream. Use as your doctor tells you. |
| Itching under the transducer arrays | Rare side effect | - Use over the counter 0.1% hydrocortisone cream when switching transducer arrays - Shift transducer arrays from the last location (so the adhesive gel is between the red marks) If the itching gets worse: 1. See your study doctor |
| Pain under the transducer arrays | Rare side effect | Stop treatment. See your study doctor. |

# ASSISTANCE & INFORMATION

Technical support

- For technical support contact your DSS. The contact information of your local DSS will be supplied to you separately.

Clinical support

- If you feel any change in your health or any side effects from the treatment call the study doctor at the center which is treating you.

# TRAVELING WITH OPTUNE®

Optune®’s batteries contain lithium ion material and are restricted from being checked as luggage for passenger aircraft travel. They can be carried in the passenger cabin.

When traveling to another country with the Optune device, use the suitable electric cable that was provided with the Optune® treatment kit.  Travel adapters should not be used with Optune® treatment kit.

Please contact your DSS if you have questions related to travel restrictions.

Note: The Optune® device and INE Transducer Arrays will activate metal detectors.

# EMITTED RADIATION & ELECTROMAGNETIC COMPATIBILITY

The Optune® treatment kit and the accompanying battery charger (ICH9100) and power supply (SPS9100) need special precautions regarding EMC and need to be installed and put into service according to the EMC information provided below.

Portable and mobile RF communications equipment can affect the Optune® treatment kit and the accompanying battery charger.

The Optune® device should be used with the following cables and additional parts only:

1. CAD9100 connection cable

2. INE9000 INE Transducer Array (Sterile)

3. IBH9100 battery

4. SPS9100 power supply

5. ICH9100 charger

6. Unshielded AC mains cables for indoor use only with a maximal length of less than 3.0 m

The use of accessories, parts and cables other than those specified, may result in increased EMISSIONS or decreased IMMUNITY of the Optune® treatment kit.

Table 1 – Guidance and MANUFACTURER’S declaration – ELECTROMAGNETIC EMISSIONS – for all ME EQUIPMENT and ME SYSTEMS

| Guidance and Manufacturer’s Declaration – Electromagnetic Emissions |  |  |
| --- | --- | --- |
| Optune Treatment Kit is intended for use in the electromagnetic environment specified below. The customer or the user of Optune Treatment Kit should ensure that it is used in such an environment. |  |  |
| Emissions test | Compliance | Electromagnetic environment – guidance |
| RF emissions CISPR 11 | Group 1 | Optune Treatment Kit uses RF energy only for its internal function. Therefore, its RF emissions are very low and are not likely to cause any interference in nearby electronic equipment. |
| RF emissions CISPR 11 | Class B | Optune Treatment Kit is suitable for use in all establishments, including domestic establishments and those directly connected to the public low-voltage power supply network that supplies buildings used for domestic purposes. |
| Harmonic emissions IEC 61000-3-2 | Class A |  |
| Voltage fluctuations/ flicker emissions IEC 61000-3-3 | Complies |  |

| Guidance and Manufacturer’s Declaration – Electromagnetic Emissions |  |  |
| --- | --- | --- |
| ICH9100 charger and SPS9100 power supply is intended for use in the electromagnetic environment specified below. The customer or the user of ICH9100 charger and SPS9100 powers supply should ensure that it is used in such an environment. |  |  |
| Emissions test | Compliance | Electromagnetic environment – guidance |
| RF emissions CISPR 11 | Group 1 | The ICH9100 charger and the SPS9100 power supply uses RF energy only for its internal function. Therefore, its RF emissions are very low and are not likely to cause any interference in nearby electronic equipment. |
| RF emissions CISPR 11 | Class B | The ICH9100 charger and the SPS9100 is suitable for use in all establishments, including domestic establishments and those directly connected to the public low-voltage power supply network that supplies buildings used for domestic purposes. |
| Harmonic emissions IEC 61000-3-2 | Class A |  |
| Voltage fluctuations/ flicker emissions IEC 61000-3-3 | Complies |  |

Warning: The Optune® treatment kit, the ICH9100 charger and the SPS9100 power supply should not be used adjacent to or stacked with other equipment.

Table 2 – Guidance and MANUFACTURER’S declaration – ELECTROMAGNETIC IMMUNITY – for all ME EQUIPMENT and ME SYSTEMS

| Guidance and Manufacturer’s Declaration – Electromagnetic Immunity |  |  |  |
| --- | --- | --- | --- |
| Optune Treatment Kit, including the Optune device, the ICH9100 charger and the SPS9100 power supply, is intended for use in the electromagnetic environment specified below. The customer or the user of Optune Treatment Kit should ensure that it is used in such an environment. |  |  |  |
| Emissions test | IEC 60601 Test level | Compliance level | Electromagnetic environment – guidance |
| Electrostatic discharge (ESD) IEC 61000-4-2 | ±6 kV contact ±8 kV air | ±6 kV contact ±8 kV air | Floors should be wood, concrete or ceramic tile. If floors are covered with synthetic material, the relative humidity should be at least 30%. |
| Electrical fast transient/burst IEC 61000-4-4 | ±2 kV for power supply lines ±1 kV for input/output lines | ±2 kV N/A |  |
| Surge IEC 61000-4-5 | ±1 kV line to line ±2 kV line to Earth | N/A |  |
| Voltage dips, short interruptions and voltage variations on power supply input lines IEC 61000-4-11 | < 5% UT (>95% dip in UT) for 0,5 cycle 40% UT (60% dip in UT) for 5 cycles 70% UT (30% dip in UT) for 25 cycles <5% UT (>95% dip in UT) for 5 s | N/A |  |
| Power frequency (50/60 Hz) magnetic field IEC 61000-4-8 | 3 A/m | 3 A/m | Power frequency magnetic fields should be at levels characteristic of a typical location in a typical commercial or hospital environment. |
| NOTE UT is the AC mains voltage prior to application of the test level. |  |  |  |

| Guidance and Manufacturer’s Declaration – Electromagnetic Immunity |  |  |  |
| --- | --- | --- | --- |
| ICH9100 charger and SPS9100 power supply are intended for use in the electromagnetic environment specified below. The customer or the user of the ICH9100 charger and SPS9100 power supply should ensure that it is used in such an environment. |  |  |  |
| Emissions test | IEC 60601 Test level | Compliance level | Electromagnetic environment – guidance |
| Electrostatic discharge (ESD) IEC 61000-4-2 | ±6 kV contact ±8 kV air | ±6 kV contact ±8 kV air | Floors should be wood, concrete or ceramic tile. If floors are covered with synthetic material, the relative humidity should be at least 30%. |
| Electrical fast transient/burst IEC 61000-4-4 | ±2 kV for power supply lines ±1 kV for input/output lines | ±2 kV N/A | Mains power quality should be that of a typical commercial or hospital environment. |
| Surge IEC 61000-4-5 | ±1 kV line to line ±2 kV line to Earth | ±1 kV line to line ±2 kV line to Earth | Mains power quality should be that of a typical commercial or hospital environment. |
| Voltage dips, short interruptions and voltage variations on power supply input lines Power frequency (50/60 Hz) magnetic field | < 5% UT (>95% dip in UT) for 0,5 cycle 40% UT 3 A/m | < 5% UT (>95% dip in UT) for 0,5 cycle 40% UT 3 A/m | Mains power quality should be that of a typical commercial or hospital environment. Power frequency magnetic fields should be at levels characteristic of a typical location in a typical commercial or hospital environment. |
| NOTE UT is the AC mains voltage prior to application of the test level = 120V and 230V |  |  |  |

Normal operation: The Optune® treatment kit is working properly when the blue LED surrounding the TTFields button are lit and no notification signal sounds. The ICH9100 charger is working properly when all the LEDs are lit. The SPS9100 power supply is working properly when the blue LEDs surrounding the TTFields button on Optune® treatment kit are lit and no notification signal sounds

Table 3 – Guidance and MANUFACTURER’S declaration – ELECTROMAGNETIC IMMUNITY – for ME EQUIPMENT and ME SYSTEMS that are not LIFE-SUPPORTING

| Guidance and Manufacturer’s Declaration – Electromagnetic Immunity |  |  |  |
| --- | --- | --- | --- |
| Optune® Treatment Kit is intended for use in the electromagnetic environment specified below. The customer or the user of Optune Treatment Kit should ensure that it is used in such an environment. |  |  |  |
| Emissions test | IEC 60601 test level | Compliance level | Electromagnetic environment – guidance |
| Conducted RF IEC 61000-4-6 Radiated RF IEC 61000-4-3 | 3 Vrms 150 kHz to 80 MHz 3 V/m 80 MHz to 2.5 GHz | 3 Vrms 3 V/m | Portable and mobile RF communications equipment should be used no closer to any part of the Optune treatment kit, including cables, than the recommended separation distance calculated from the equation applicable to the frequency of the transmitter. Recommended separation distance 80 MHz to 800 MHz 800 MHz to 2.5 GHz Where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer and d is the recommended separation distance in meters (m). Field strengths from fixed RF transmitters, as deter-mined by an electromagnetic site surveya, should be less than the compliance level in each frequency range. b) Interference may occur in the vicinity of equipment marked with the following symbol: |
| NOTE 1 At 80 MHz and 800 MHz, the higher frequency range applies. NOTE 2 These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. |  |  |  |
| - Field strengths from fixed transmitters, such as base stations for radio (cellular/cordless) telephones and land mobile radios, amateur radio, AM and FM radio broadcast and TV broadcast cannot be predicted theoretically with accuracy. To assess the electromagnetic environment due to fixed RF transmitters, an electromagnetic site survey should be considered. If the measured field strength in the location in which Optune Treatment Kit is used exceeds the applicable RF compliance level above, Optune Treatment Kit should be observed to verify normal operation. If abnormal performance is observed, additional measures may be necessary, such as re-orienting or relocating Optune Treatment Kit - Over the frequency range 150 kHz to 80 MHz, field strengths should be less than 3 V/m |  |  |  |

Table 4 – Recommended separation distances between portable and mobile RF communications equipment and the ME EQUIPMENT or ME SYSTEM – for ME EQUIPMENT and ME SYSTEMS that are not LIFE-SUPPORTING

| Optune Treatment Kit is intended for use in an electromagnetic environment in which radiated RF disturbances are controlled. The customer or the user of Optune Treatment Kit can help prevent electromagnetic interference by maintaining a minimum distance between portable and mobile RF communications equipment (transmitters) and Optune Treatment Kit as recommended below, according to the maximum output power of the communications equipment. |  |  |  |
| --- | --- | --- | --- |
| Rated maximum output power of transmitter W | Separation distance according to frequency of transmitter m |  |  |
|  | 150 kHz to 80 MHz | 80 MHz to 800 MHz | 800 MHz to 5.5 GHz |
| 0.01 | 0.116 | 0.116 | 0.233 |
| 0.1 | 0.368 | 0.368 | 0.736 |
| 1 | 1.16 | 1.16 | 2.33 |
| 10 | 3.68 | 3.68 | 7.36 |
| 100 | 11.6 | 11.6 | 23.2 |
| For transmitters rated at a maximum output power not listed above, the recommended separation distance d in metres (m) can be estimated using the equation applicable to the frequency of the transmitter, where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer. NOTE 1 At 80 MHz and 800 MHz, the separation distance for the higher frequency range applies. NOTE 2 These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. |  |  |  |