Back-translation

®

Optune® is a CE-marked device.

For investigational use only.

| Study | EF-41 (KEYNOTE D58) |
| --- | --- |
| Sponsor | Novocure GmbH Park 6 CH-6039 Root D4 Switzerland |
| Document number | PUM-SD-670 ES BT |
| Version: | 1.0 |

Table of contents

1.	GLOSARIO DE TÉRMINOS MÉDICOS	3

2.	USO PREVISTO	4

3.	¿QUÉ ES OPTUNE®?	4

4.	CONTRAINDICACIONES, ADVERTENCIAS Y PRECAUCIONES	5

5.	 ACERCA DEL KIT DE TRATAMIENTO OPTUNE®	9

6.	EL DISPOSITIVO OPTUNE®	10

7.	LOS INE TRANSDUCER ARRAYS	12

8.	ANTES DE EMPEZAR	12

9.	EXTRACCIÓN DE LOS INE TRANSDUCER ARRAYS DE SU ENVOLTORIO	13

10.	PREPARACIÓN DE LA CABEZA PARA LA COLOCACIÓN DE LOS INE TRANSDUCER ARRAYS	14

11.	COLOCACIÓN DE LOS TRANSDUCTORES EN LA CABEZA	16

12.	CONEXIÓN DE LOS INE TRANSDUCER ARRAYS AL DISPOSITIVO	18

13.	EL CABLE DE CONEXIÓN	19

14.	INICIAR Y DETENER EL DISPOSITIVO	20

15.	CONEXIÓN Y DESCONEXIÓN DE LA BATERÍA	25

16.	CARGAR LA BATERÍA	29

17.	USO DEL TRANSFORMADOR DE ALIMENTACIÓN ENCHUFABLE	32

18.	DESCONEXIÓN DEL DISPOSITIVO	33

19.	TRANSPORTE DEL DISPOSITIVO	35

20.	GLOSARIO DE SÍMBOLOS GRÁFICOS	36

21.	CONDICIONES AMBIENTALES PARA EL FUNCIONAMIENTO, EL ALMACENAMIENTO Y EL TRANSPORTE	39

22.	RESOLUCIÓN DE PROBLEMAS	40

23.	ASISTENCIA E INFORMACIÓN	43

24.	VIAJAR CON OPTUNE®	44

25.	EMISIÓN DE RADIACIÓN Y COMPATIBILIDAD ELECTROMAGNÉTICA	45

# GLOSSARY OF MEDICALS TERMS

Cancer – abnormal cell division that propagates uncontrollably

Chemotherapy – drugs that should destroy or damage the cancer cells

Clinical study – research study in which people participate

Contraindications – situations in which a treatment should not be used

Device support specialist (DSS) – Novocure device technician who provides help with the technical aspects of Optune®

Electric field generator (the device) – portable device to administer TTFields to patients’ brain

Glioblastoma (GBM) – a type of brain cancer

INE Transducer Array (also called transducer array) – adhesive bandages that hold the insulated ceramic discs that administer TTFields to the scalp.

Local – in one part of the body

Pembrolizumab – drug used to treat cancer

Radiation – treatment with X-rays used to eliminate tumor cells

Corticosteroid – when administered orally or by IV (in a vein), medication that is used to reduce the swelling that surrounds a brain tumor and helps with brain symptoms. When used on skin, medication that can reduce inflammation

Temozolomide (TMZ) – type of cancer medication used to treat GBM

Topical – on the skin surface

TTFields – (Tumor Treating Fields): alternating electrical fields, administered by transducer arrays to the part of the body where the solid tumor is located. TTFields have been shown to destroy tumor cells.

Tumor – abnormal tissue growth

# INTENDED USE

In this clinical study, the use of Optune® concomitant with temozolomide and pembrolizumab/placebo is intended for the treatment of newly diagnosed supratentorial glioblastoma after maximal cytoreductive surgery and completion of treatment with radiation therapy (RT) concomitant with standard of care chemotherapy.

# WHAT IS OPTUNE®?

Optune® is a portable battery-powered device that administers electric fields, called tumor treating fields (TTFields) to the patient through insulated INE Transducer Arrays.

TTFields produce electric forces with the purpose of destroying cancer cells in the brain. The device and battery are carried in a shoulder bag or on the shoulder of the opposite site.

Treatment with Optune® is intended to be used with a minimum monthly average use of 18 hours daily.

# CONTRAINDICATIONS, WARNINGS AND PRECAUTIONS

Contraindications

Do not use Optune® if you are under 18 years of age.

Do not use Optune® if you have an implanted electronic medical device, a defect in the skull (such as missing, unreplaced bone) or bullet fragments. Examples of active electronic devices are deep brain stimulators, spinal cord stimulators, vagus nerve stimulators, pacemakers, defibrillators and programmable leads. The use of Optune® with implanted electronic devices has not been tested, hence, it may cause malfunction of the implanted device. The use of Optune® in persons with defects or bullet fragments in the skull has not been tested, hence, there is a possibility of tissue damage or of Optune® being ineffective.

Do not use Optune® treatment kit if you know that you are sensitive to the conductive hydrogels, such as the gel used on the electrocardiogram (ECG) adhesive pads or on transcutaneous electrical nerve stimulation (TENS) electrodes. In this case, skin contact with the gel used with the Optune® treatment kit may cause an increase of redness and itching and, in rare cases, may even can cause severe allergic reactions such as anaphylactic shock and respiratory failure.

Do not use the Optune® treatment kit if you are pregnant or breastfeeding.

Warnings

Warning: use Optune® only after receiving training from qualified personnel.

Warning: in case of skin irritation, which manifests in the form of redness underneath the INE Transducer Arrays (mild rash), use a high-potency topical corticosteroid when you replace the transducer arrays. This will help relieve the skin irritation. If you do not use a cream of this type, the skin irritation may become more severe and may even cause the appearance of ruptures, infections, pain and blisters in the skin. If so, stop using the topical corticosteroid cream and contact your doctor. He/she will provide an antibiotic cream for you to use when you replace the transducer arrays. If you do not use it, the symptoms may continue and the doctor might ask you to take a break in the treatment until the skin heals. Taking a treatment break may reduce your likelihood of responding to treatment.

Warning: all repair procedures must be performed by qualified and properly trained personnel. If you try to open the device yourself, you can damage it. You might also receive an electrical discharge when you touch the internal parts of the device.

Precautions

Caution: keep Optune® out of the reach of children. If children touch the device, they may damage it. This could cause a pause in the treatment. Pauses in the treatment may reduce the probability of responding to it.

Caution: do not use any part that does not come with the Optune® treatment kit or that has not been supplied by the device manufacturer or given to you by your doctor. The use of other parts manufactured by other companies or for use with other devices may damage the device. This could cause a pause in the treatment. Pauses in the treatment may reduce the probability of responding to it.

Caution: if your doctor used plates or screws to close the skull bone during the surgery, be careful when placing the transducer arrays.  Make sure the round discs that make up the transducer arrays are not over any area where you can feel the screws or plates underneath the skin.  Otherwise, make sure that the screws or plates you have underneath your skin are between the round discs that make up the transducer arrays.  If you do not do so, it may increase skin damage, which may lead to a break in the treatment.  Pauses in the treatment may reduce the probability of responding to it.

Caution: tell your doctor before using the device if you have any inactive implanted medical device in your brain (such as a stent, a plastic reservoir for drug delivery, an aneurysm clamp or coil or a device lead).  The use of Optune® has not been tested in patients with inactive implanted medical devices in the brain has not been tested, since they could cause tissue damage or reduce your chances of response to the treatment.

Caution: do not use Optune® if any part appears to be damaged (torn cables, loose connectors, loose sockets, cracks or breaks in the plastic housing). The use of damaged components can deteriorate the device and lead to a treatment break. Pauses in the treatment may reduce the probability of responding to it.

Caution: do not wet the device or the transducer arrays. Wetting the device may damage it, which would prevent you from receiving treatment for enough time. If the transducer arrays get very wet, they are likely to detach from the head, in which case the device will turn off and you will have to change the transducer arrays.

Caution: before connecting or disconnecting the transducer arrays, make sure that the Optune® power switch is in the “OFF” position. Disconnecting the transducer arrays with the device power switch in the “ON” position could trigger the alarm sounds and may damage the device.

Caution: if you have a severe underlying skin condition on the scalp, talk to your doctor to find out if this could hinder treatment with Optune® or temporarily interfere it.

Notices

Notice: The Optune® device and the transducer arrays activate metal detectors.

Notice: You should use Optune® for at least 18 hours a day to obtain the best response to the treatment. The use of Optune® less than 18 hours a day reduces the possibilities of responding to the treatment.

Notice: Do not stop using Optune® even if you have used it less than the recommended 18 hours a day. You should only stop using the device if you are told to do so by your doctor. Stopping the treatment could reduce the possibilities of responding to the treatment.

Notice: If you plan to be away from home for more than 2 hours, take an extra battery and/or the power supply with you in case the battery you are using runs out. If you do not take a spare battery or the power supply, you may need to pause your treatment. Pauses in the treatment may reduce the probability of responding to it.

Notice: Make sure that you have at least 12 spare transducer arrays at all times. That way, you will have enough until you receive the next shipment of transducer arrays. Remember to request more transducer arrays when you have at least 12 spare transducer arrays. If you do not order transducer arrays in time, a break in treatment may occur. Pauses in the treatment may reduce the probability of responding to it.

Notice: Batteries may weaken over time and have to be replaced. You will know this is the case when the amount of time the device runs on a fully charged battery begins to shorten. For example, if the low battery indicator flashes only an hour and a half after starting treatment, replace the battery. If you have no spare batteries when the batteries run out, there may be a pause in treatment. Pauses in the treatment may reduce the probability of responding to it.

Notice: You should have the troubleshooting guide with you (section‎22) at all times.  This guide is necessary to ensure that the Optune® operates correctly.  If you do not use the device properly, a break in treatment may occur.  Pauses in the treatment may reduce the probability of responding to it.

Notice: Do not block the device vents located at the front and back of the electric field generator. Obstructing the vents may make the device overheat and turn off, causing a pause in the treatment. If this happens, unobstruct the vents, wait 5 minutes and restart the device.

Notice: Do not obstruct the vents of the battery charger located on its side. Obstructing the vents may make the charger overheat, which could prevent the batteries from charging.

Notice: Before using a transducer array, make sure that the package is sealed by rubbing it gently between your thumb and index finger on all four sides. The package must be closed on all sides. There must be no openings in the package seal. If the package is not sealed, the transducer arrays could be damaged. Damaged transducer arrays will not work properly, which could make the device turn off.

Notice: The transducer arrays are for single use and should not be removed from the head and put back on. If you place a used transducer array back on your head, it may not adhere well to the skin and the device may turn off.

# ABOUT THE OPTUNE® TREATMENT KIT

| Optune® treatment kit | Model number |
| --- | --- |
| 1 Electric field generator (the device) | TFH9200 |
| 2 Battery | IBH9100 |
| 3 Battery charger | ICH9100 |
| 4 Plug-in power supply | SPS9100 |
| 5 Connection cable and box (CAD) | CAD9100 |
| 6 INE Transducer Arrays | INE9TAN and INE9TANW |
| 7 Power cables | N/A |
| 8 Shoulder bag and strap | N/A |
| 9 Battery case | N/A |

# THE OPTUNE DEVICE

Optune® is a portable medical device. It administers electric fields called “TTFields” to the brain through INE Transducer Arrays.  Maintain the TTFields treatment at all times, as long as possible, when you are awake and when you are sleeping.  Limit treatment interruptions to the shortest time possible.  The image below shows the controls of the device for operation of the system.

It is not necessary to adjust any settings. You only need to turn the device and the therapy on and off.

The device has the following controls:

- Power switch

- Power supply port

- TTFields therapy ON/OFF button

- POWER / BATTERY / ERROR indicators

- Connection cable (CAD) port

- Battery gauge

- Battery charge indicator button

# INE TRANSDUCER ARRAYS

The INE Transducer Arrays are transducer arrays with insulation that administer tumor treating fields (TTFields) to a tumor in a non-invasive manner.

The INE Transducer Arrays are supplied sterile and should only be used with the Optune® treatment kit.  The INE Transducer Arrays do not require maintenance, cleansing or disinfection.

# BEFORE STARTING

You will need to use four (4) transducer arrays (two black and two white) each time you replace the transducer arrays. The black transducer arrays are placed on the front and back of the head, while the white transducer arrays are placed on the sides of the head. Remember: the black ones go on the back and the white ones go on the right. Change the four (4) transducer arrays at least twice a week (at most, every 4 days) to continue the treatment with Optune®.

If necessary, a doctor, nurse or caregiver can help you change the transducer arrays.

# REMOVING THE INE TRANSDUCER ARRAYS FROM THEIR WRAPPING

Open the clear package of the four (4) INE Transducer Arrays by gently separating the opposite edges of the envelope. Hold the INE Transducer Array as shown in the image.

# PREPARING THE HEAD FOR PLACEMENT OF THE INE TRANSDUCER ARRAYS

Wash your hands thoroughly and wash your head with a gentle shampoo.

If this is the first time you are using the transducer arrays, skip this step and go directly to the next step (shaving).  If you are replacing the transducer arrays, you, or your doctor or caregiver if necessary, should clean the skin with baby oil to remove any adhesive from previous transducer arrays. This will not prevent the correct operation of the device.

Shave your entire scalp with an electric razor. Do not leave any trace of hair.

Clean your scalp with 70% alcohol.  If your skin is irritated, avoid applying alcohol and wash your head with a gentle shampoo and water before drying it gently with a soft towel.

If your scalp is red, apply the corticosteroid cream prescribed to you by your doctor.

If you have any sores on your scalp, treat them as instructed by your treating physician.

Wait at least 15 minutes and then gently clean your skin again with 70% alcohol to remove any residue.

Once the scalp is dry, apply the transducer arrays.

# PLACING THE TRANSDUCER ARRAYS ON YOUR HEAD

After preparing your scalp, place the INE Transducer Arrays on your head with help of a doctor or caregiver if necessary. At least twice a week (every 4 days, at most), remove the INE Transducer Arrays, prepare your scalp and place a new set of transducer arrays.

You will know it is time to change the transducer arrays when the device emits a beep more frequently. This means that the device cannot work properly due to hair growth. Hair growth prevents the transducer arrays from establishing a good contact with the scalp.

To place the INE Transducer Arrays on your head, with help from a caregiver or doctor if necessary, follow the steps below.

Note: if this is the first time you are using the INE Transducer Array, skip the first step (removal).

  - Remove the transducer arrays from your head by pulling the medical tape from the scalp.

  - Check what color transducer array goes on what part of the head. The black transducer arrays are placed on the front and back of the head, while the white transducer arrays are placed on the sides.

  - Prepare the skin for the transducer arrays, as explained above.

  - Peel off the white covering that covers the gel of the first transducer array.

If this is the first time that you are using the transducer arrays, place them on your head as shown in the transducer arrays placement diagram (“map”) given to you by your doctor.

The placement points depend on where your tumor is located. When you replace the transducer arrays, place them on your head in the same location as before, but shift them less than an inch in the direction of the arrow on the transducer array placement diagram (“map”). To reduce skin irritation under the transducer arrays, move them a bit. It is not necessary to move the transducer arrays for the device to work properly.

The ceramic gel discs leave a slight impression on the surface of the scalp. This will help you place the new transducer arrays. The transducer arrays act in pairs, so they should always be moved together (see the figure). The next time the transducer arrays are replaced, move them back to the previous location.

Patient 1	Patient 2	Patient 3

Patient 1	Patient 2	Patient 3

Place the other three transducer arrays in the same way.

Pull from the tabs on each side of the transducer arrays and press them firmly against your scalp.

You may need the help of a friend or family member to place the back transducer arrays.

Press the entire edge of the INE Transducer Array tape against the skin.

# CONNECTING THE INE TRANSDUCER ARRAYS TO THE DEVICE

Connect the four black and white connectors of the INE Transducer Arrays to the respective sockets coded in black and white on the Optune® connection cable. Plug the transducer array connectors with the black ends into the two black sockets (one of them will be labeled “P1” and the other “N1”) and the two white ends into the white sockets (one of them will be labeled “P2” and the other “N2”) (see the diagram).

Press firmly to ensure that the connectors are fully inserted.

Gather the INE Transducer Array cables together and secure them at a place that is comfortable for you, using a small piece of tape. You can secure the clip of the connection cable to your belt.

# CONNECTION CABLE

The connection cable is the wound-up flexible cable that goes from the device to the connection box. The four black (2) and white (2) connectors of the transducer arrays are plugged into the connection box. The black and white coding matches the position of the transducer arrays on the head.

Plug the connection cable into the socket at the left of the front panel of the device. The connection cable is plugged into the socket with the connector arrow pointing upward.

Push the connector until you hear a click. The click means it is connected.

# STARTING AND STOPPING THE DEVICE

To turn on the device and start the treatment:

Place the transducer arrays on your scalp (if necessary, with the help of a caregiver) as described in section ‎0 above.

Connect the transducer arrays to the connection cable box, as described in section ‎12 above.

Plug the connection cable into the connector of the device with the arrow up (as described in section ‎13 above).

Insert a charged battery into the device (see section ‎15) or connect the wall power supply (section ‎17).

Place the power button in the “ON” position.  The power button is located on the bottom of the device.

Wait about 10 seconds for the self-check to be completed. The “POWER” indicator of the front panel of the device will light up in green.

If a charged battery has been installed and the device is not plugged into a wall power supply, the “BATTERY” indicator will also light up in green.

If the device is plugged into a wall power supply connected, the device will operate with the wall power supply and the “BATTERY” indicator will not light up.

Press the TTFields therapy button once; treatment will then start.

The word “TTFIELDS” located above the TTFields therapy button will light up in blue and remain on while the treatment is active. If the blue word does not light up in blue, the treatment is not functioning, so you should check the settings and restart the procedure.

If the indicators do not turn on after doing so, refer to the troubleshooting guide (section ‎22). If you continue having problems, contact technical support or your DSS (section ‎23).

Note: The green, blue and yellow indicators will dim in a dark room and shine more brightly in a lighted environment. The red “ERROR” Indicator will not dim in any case.

If you do not press the therapy button in the 10 minutes after turning on the device, a notification signal will sound that indicates that the device is “ON” but the therapy is “OFF.” This is a reminder to start the therapy. To silence the alarm, you should press the therapy button once, and again to start the therapy. “TTFIELDS” will light up in blue when the patient is receiving treatment.

Turning off the device and stopping the treatment

You can stop the device when the following happens:

- If the device is operating correctly, but you need to stop the treatment to take a break:

Press the TTFields button. The blue “TTFIELDS” lights around the button will turn off, which stops the TTFields therapy, but the device remains on.

Then, turn off the device by setting the power button at the bottom of the device to the “OFF” position.

- If an error occurs:

If an error occurs, the device will turn off the TTFields and will issue a loud beep.

The red “ERROR” light will come on (as shown below). In that case, once the red “ERROR” light comes on, the treatment will have stopped.

To turn off the device:

- Press the TTFields button at the front of the device to stop the alarm. The red “ERROR” light will turn off. (If this does not silence the alarm, go to the next step and turn off the power switch to silence it. There are a few alarm conditions that cannot be silenced by pressing the TTFields button).

- Turn off the device by setting the power button to the “OFF” position.

- To obtain instructions on how to solve problems, refer to the troubleshooting guide (section ‎22).

If you do not find any problems, restart the device and restart the treatment. If the alarm continues, contact technical support or the DSS (section ‎23).

- If the low battery indicator turns on:

When the battery runs out (after about two or three hours), an alarm beep will sound, the TTFields therapy will stop and the yellow low “BATTERY” light and the red “ERROR” light will come on. The sound of this alarm is the same sound emitted by the device to indicate that an error has occurred. Nevertheless, in this case both the yellow light and the red light will light up, instead of only the red light.

To turn off the device:

- Press the TTFields button at the front of the device to stop the alarm. The red “ERROR” and yellow low “BATTERY” lights will turn off.

- Replace the battery by following the steps in section ‎15.

# CONNECTING AND DISCONNECTING THE BATTERY

The Optune® treatment kit is supplied with 4 rechargeable batteries. The batteries slide into the device. The battery must be inserted until you hear a “click,” which indicates that the battery is in place. Take care to avoid letting the battery fall into place or forcing it into the battery slot. The Optune® device uses one battery at a time. The other three batteries should remain in the battery charger. Every battery lasts from 2 to 3 hours. Replace the battery whenever it is depleted (when the yellow low “BATTERY” indicator turns on, as explained in section ‎14). If you plan to be away from home for more than two hours, take spare batteries or a power supply with you.

Press gently downward to latch the battery into place. Make sure that the battery pin is completely coupled.

To remove the battery from the slot, press both blue buttons at the sides of the battery and slide it upward.

Recharge the batteries in the charger (see section ‎8) for two to four hours. The batteries will keep most of their charge for several days after they are removed from the charger, but will eventually lose it all.

Keeping the batteries in the charger once they are fully charged will not damage them, so you can leave them there if you do not need them.

You can charge and use the batteries many times for about six to nine months.

Over time, the period during which the batteries can power the device (before the yellow low “BATTERY” indicator lights up and the alarm sounds) will shorten.

If the time from starting the treatment with the full battery until the low battery alarm sounds is below 90 minutes, contact technical support or the DSS (see section ‎23) to obtain spare batteries.

The “BATTERY” light will switch from green to yellow when the battery charge falls below a threshold, which indicates that the battery should be changed soon.

The treatment will continue while the yellow low “BATTERY” indicator is lit until the audible alarm sounds and the red “ERROR” light lights up.

When this happens, the treatment stops and the device should be turned off and the battery replaced.

When the yellow low “BATTERY” indicator turns on, there are two ways to replace the depleted battery with a charged battery.

Option one: (use it if you are near the direct wall power supply) allows you to change the battery without interrupting the therapy. You may do so before the battery is fully depleted and before the device activates the notification signal.

Follow the next steps:

- Connect the wall power supply to the back of the Optune® device. (see section ‎17). The lights on the display panel will indicate that it is no longer using battery power.

- Remove the battery from the battery slot by pressing the blue buttons on both sides of the battery and remove the battery cartridge by lifting it out of the slot.

- Select a fully charged battery from the charger or device bag.

- Slide the fully charged battery into the battery slot and press it gently downward to fit it in place.

- Remove the power cord from the bottom of the device.

The “BATTERY” light on the display panel will light up, which indicates that it is no longer functioning with the wall power supply.

To check the battery gauge, see the next section.

Option two: if you are not near the wall power supply, or if the battery has been completely depleted, replace the battery following these steps:

- Press the TTFields button to stop the treatment.

- Turn off the device with the power switch (at the bottom of the device).

- Remove the battery from the battery slot by pressing the blue buttons on both sides of the battery and remove the battery cartridge by lifting it out of the slot.

- Select a fully charged battery from the charger.

- Slide the fully charged battery into the battery slot and press it gently downward to fit it in place.

- To check the battery gauge, see the next section.

- Put the device power switch in the “ON” position and wait for the self-check to complete (it takes about 10 seconds). To start the treatment, press the TTFields button.

- Insert the used battery into the battery charger to recharge it (see section ‎8).

# CHARGING THE BATTERY

Checking the battery gauge

While using the Optune® treatment kit, you may want to check how much power is left the battery. Checking the battery will not interfere with your treatment or stop it. To check how much charge is left in battery, press once on the button at the top of the battery pack. The battery charge will be displayed on the lighted meter to the right of the button. The indicator goes from full charge to empty, like a fuel gauge in a car.

Charge full

75% charge

50% charge

25% charge

Empty

Charge full

75% charge

50% charge

25% charge

Empty

Charging the batteries

The battery charger recharges used batteries. To do so, it uses power from a standard wall socket. Each battery fits into a slot that connects it directly to the charger.

Before charging the batteries, plug in the charger power cord to a standard wall socket and turn the power button on at the back of the charger.

The front lights of the charger will turn on during the self-check and then the small light in the center of the front panel will light up in green to indicate that power is being applied.

To recharge a used battery:

- Place the used battery in one of the three openings on top of the charger. Slide the battery until it is fully fitted in place.

- The light located directly in front of the opening where the battery is plugged in will flash green, indicating that it is charging. The green light will flash more quickly once the battery has charged up to 95% of its capacity. You can also check the battery gauge during charging to obtain information on the charge level in the battery.

- When the battery is fully charged (around 2 to 4 hours), the charge light will go from flashing green to steady green. The steady green light will disappear when the battery is extracted or the charger is disconnected from the standard wall socket.

- Charger power cord

- Power switch

- Charger power cord

- Power switch

Rear view of the battery charger that shows the power switch

and where the power cord is connected

If the light in the front panel turns red, this indicates that there is a problem with the battery or charger, hence, you should contact technical support or the DSS to get help. Do not use batteries that cause a red light to turn on in the charger.

Keep the batteries in the charger even after they are fully charged. This will not damage the batteries.

- Battery charging slot

- Battery charge indicator

- Charger indicator

- Battery charging slot

- Battery charge indicator

- Charger indicator

Front view of the battery charger showing how the batteries are inserted into the charger

Notice: The charger is not intended to be used in the presence of flammable mixtures.

# USING THE PLUG-IN POWER SUPPLY

If you intend to stay in one place for a while, such as when you are sleeping, you can use the plug-in power supply instead of the batteries. Unlike the batteries, there is no limit to how long the device can function when the plug-in power supply is used. The plug-in power supply functions with U. S. (120 VAC) or European (230 VAC) power outlets.

Note: It is normal for the power supply to become warm when it is being used. If the power supply is too hot to touch, unplug it and contact technical support or the DSS (section ‎23).

When the device has a battery in place and is also connected to the wall power supply, the device uses the latter as the preferential power source. When the wall power cord is plugged in while the device is running on the battery, the device automatically switches from battery power to power from the wall power supply.

### Connecting the plug-in power supply

- Connect the power supply to a standard wall socket with the power cord supplied with it.

- To use the wall power supply, it is not necessary to remove the battery from the device. Keep in mind that the device battery will not charge when the device is connected to the wall power supply. Depleted batteries must be placed in the battery charger to recharge them. If the TTFields are activated, is not necessary to turn them off to plug in the wall power supply.

- Plug the wall power supply connector into the AC port on the back of the device (next to the power switch).

- If TTFields are functioning, the device will automatically switch to the wall power supply without interrupting the TTFields. If the device is not on, turn on the power switch and wait until the self-check has been completed (about 10 seconds). Press the TTFields button to start the device (as explained in section ‎14).

### Disconnecting the plug-in power supply and returning to battery power

- Make sure that a charged battery is properly inserted into the device before removing the wall power supply. If TTFields are functioning, it is not necessary to turn them off before removing the wall power supply. Once the power supply is removed, the device will automatically switch to battery power.

- Remove the connector from the plug-in power supply at the back of the device.

- If the device is not on, turn on the power switch and wait until the self-check has been completed (about 10 seconds). Press the TTFields button to start the device.

- Keep the plug-in power supply for the next time you use it.

# DISCONNECTING THE DEVICE

There are two ways to disconnect the device to pause the treatment (after turning off the device)

- Disconnect the connection cable from the device.

- Disconnect the transducer arrays from the connection cable.

Unplugging the connection cable from the device

- Stop the TTFields treatment by pressing the TTFields button.

- Turn off the Optune® device with the power switch.

- Disconnect the connection cable from the socket by holding the connector and pulling.

PRECAUTION! Do not pull on the cable!

Now you can move about without the device, but will still be connected to the connection cable and box.

To start treatment again after the break:

- Plug the connection cable into the device socket with the arrows pointing upward.

- Turn on the device with the power switch. Wait until the self-check is completed (about 10 seconds).

- Turn on the TTFields with the TTFields button.

To disconnect the transducer arrays from the connection cable box

To take a break from the treatment and disconnect completely from the device, disconnect the transducer array cables from the connection cable box.  The connection cable remains plugged into the device socket.

- Stop treatment by pressing the TTFields button.

- Turn off the Optune® device with the power switch.

- Disconnect the connectors of the transducer arrays from the connection box by pulling on the connectors.

Note: You may need to move the transducer array cables to remove them.

To restart the treatment, plug the transducer arrays into the connection box. Connect each transducer array to the respective color (black or white) according to the position of the transducer arrays on your head (as described in section ‎0).

- Once the 4 transducer arrays are connected, turn on the power switch of the device and wait until the self-check is completed (about 10 seconds). Press the TTFields button to restart the treatment.

# CARRYING THE DEVICE

Both the electric field generator and the battery fit in a carrying bag. The bag can be carried two ways: by the handle, with the supplied shoulder strap, or on the shoulder of the opposite site.

Note: Do not place the device in a different bag. The Optune® device has a fan inside that needs air flow. The bag supplied with the device is designed to allow adequate air flow. If you place the device in a bag without adequate airflow, it could overheat and the treatment would stop. If this happens, you will hear an alarm.

# GLOSSARY OF GRAPHIC SYMBOLS

|  | Precaution Check the instructions for use to see important precaution information, such as warnings and precautions |
| --- | --- |
|  | Date of manufacture |
|  | Fragile: handle with care |
|  | Follow the instructions for use |
|  | Keep dry. Do not enter rooms with high humidity or danger of direct exposure to water while wearing the device. Do not use the device if it is not in the carrying bag. Do not expose the device to direct rain. |
|  | For indoor use only The charger and the power supply are for indoor use only. |
|  | Recycling of waste from electrical and electronic equipment (WEEE). Contact technical support to arrange for proper disposal of used transducer arrays or transducer arrays that are no longer used. |
|  | The batteries are lithium ion. Contact technical support to arrange for proper disposal of dead batteries or batteries that are no longer used. |
|  | Protect from heat and from radiation sources Optune® must be keep away from extreme heat and radiation sources. |
|  | Type BF applied part: symbolizes the part that comes into contact with the patient. Specifies the P/N of the applied part that will be used with this device. |
|  | Expiry date: do not use after this date |
|  | ON/OFF switch for the electric field generator and the battery charger: When the switch is in the I position, the device is “ON.” When the switch is in the O position, the device is “OFF.” |
|  | Do not use if the package is damaged Do not use the transducer arrays if the packaging is damaged. |
|  | Do not reuse. The INE Transducer Arrays are for single use and must not be reused. |
|  | Do not resterilize |
|  | Sterilized/sterilization method The INE Transducer Arrays are sterilized using γ radiation. |
|  | Model number |
|  | Catalog number |
|  | Serial number |
|  | Lot number |
|  | Class II equipment as per IEC 60601-1 |
| IP21 | Protects people from reaching hazardous parts with their fingers. Protects the equipment inside the housing from the entry of solid foreign objects with a diameter of 12.5 mm or more. Protects the equipment inside the housing from the entry of water drops falling vertically. |
| IP22 | Protects people from reaching hazardous parts with their fingers. Protects the equipment inside the housing from the entry of solid foreign objects with a diameter of 12.5 mm or more. Protects the equipment inside the housing from the entry of drops of water that fall vertically when the housing has a maximum tilt of 15°. |
|  | Storage temperature range Do not expose them to temperatures below –5°C or above 40°C |
|  | Storage humidity range Do not expose to humidity under 15% or above 93% |
|  | P1 P2 N1 N2 codes in black and white on the connection box |
|  | Manufacturer information: Novocure GmbH, Business Village D4, Park 6/Platz 10, 6039 Root, Switzerland |

# ENVIRONMENTAL CONDITIONS FOR OPERATION, STORAGE AND TRANSPORT

Operating conditions

All treatment kit components will be normally used under the conditions specified below:

- Mainly for home use

- For indoor use only (charger and power supply)

- Not suitable for the shower, bath or bathroom, or to go out under heavy rain

- Do not use in the presence of flammable mixtures

- It can be allowed to fall on the floor without safety risks to the patient, but it may stop working

Conditions of visibility:  none.

Cleaning

All treatment kit components can be periodically cleaned with a damp cloth to remove dust and usual dirt.

Physical operating conditions for all treatment kit components:

- Temperature range: -5 °C/40 °C (23 °F/104 °F)

- Relative humidity range: 15%-93%

- Environmental pressure range: 700 to 1060 hPa

Storage conditions

- Temperature range: -5 °C/40 °C (23 °F/104 °F) for the device and additional parts

- Temperature range: 5 °C/27 °C (41 °F/80 °F) for the INE Transducer Arrays

Transport conditions

The device, additional parts and INE Transducer Arrays can be transported by air or ground, provided that this is done under weather-protected conditions, as described below:

- Temperature range: -5 °C/40 °C (23 °F/104 °F)

- Maximum relative humidity: 15%-93%

- Do not expose directly to water.

# TROUBLESHOOTING

| Problem | Possible causes | Measures to be taken |
| --- | --- | --- |
| The device power indicator does not come on after placing the switch in the “ON” position | - Device not connected to the power source - If it works with the battery: battery is depleted - Battery malfunction - If it works with the power supply: not properly plugged into the wall - The device is not working correctly - Power supply malfunction | - If it is running on the battery: check the battery gauge to confirm that it is not depleted. If so, replace it with a fully charged battery or connect it to the power supply. - Check that both the device and the power source are properly connected and try again - Assess the integrity of all connectors. None of the elements should be damaged or broken in any way - If the device cannot be turned on with the battery or the wall power supply, or if a part is damaged, do not use the device - Contact the DSS for technical support |
| A cable has become disconnected from the transducer arrays / the connection cable / the device | - Excessive physical force on the cables - The device is not working correctly | - Silence the notification signal by pressing the TTFields button - Assess the connectors. If they are intact, reconnect and restart the therapy - If something looks damaged or cannot be connected properly, do not attempt to use the device - Contact the DSS for technical support |
| The device has fallen or gotten wet | Incorrect use | - Press the TTFields button to stop the therapy - Turn OFF the power switch - Contact the DSS for technical support |
| The device alarm is activated | - Low battery - Loose or disconnected cable - The device is too hot: the device air vents are blocked - Localized hotspot on the transducer array from being placed on a pillow, for example - Poor contact of the transducer array due to hair growth or another reason - The device is not working correctly - The device is “ON” but the therapy has not been activated | If the low battery light is yellow: - Replace the battery as explained above in section ‎8 - Turn on the treatment If the error light comes on, but the low battery light is green or is not on: - Press the TTFields button to stop the alarm - Wait several seconds, then press the TTFields button again - If the blue lights around TTFields button turn on, the therapy has already been activated If the notification signal repeats after a few minutes: - Silence the notification signal and turn the device completely off - Disconnect all the equipment and make sure that nothing is loose, damaged or broken. If that is not the case, replace the damaged component before trying to turn the device on again - Reconnect all equipment in the correct order and turn the device on again. Check that the self-check is completed and press the TTFields button - Check the device vents to make sure that they are not blocked - If you are lying down, move your head - Make sure that the transducer arrays are properly adhered to your head; add adhesive tape if necessary - Restart the treatment - If the alarm continues, turn off the device and contact the DSS |
| The notification alarm sounds 10 minutes after you turn on the device. | 1. Standby therapy time | The device notification alarm will sound if it is turned on for 10 minutes, but therapy is not started. This is a reminder to start the therapy and does not indicate a malfunction. - Silence the notification alarm by pressing the TTFields button, then wait several seconds and press the TTFields button again to start the treatment. The blue indicator around the TTFields button will turn on to indicate that the therapy is now activated - If other alarms sound, review the general “notification signal” section below |
| The low battery indicator remains on after replacing the battery | - The charger is not working correctly - Battery malfunction - The device is not working correctly | - Replace the battery with a spare fully charged battery - If the problem is not resolved, contact the DSS for technical support |
| When the device is turned on a continuous notification alarm sounds and all the lights remain on indefinitely. The device does not complete the self-check | - The device is too hot - The device is not working correctly - The power source is not working correctly | - Turn off the device completely with the power switch - Check that the device is not hot to touch - Connect the device to a different power source and try to turn it on again - If the device cannot be turned on with the battery or the wall power supply, or if a part is damaged, contact DSS for technical support |
| Reddening of the skin under the transducer arrays | Common side effect | When you change the transducer arrays, use an over-the-counter 0.1% hydrocortisone cream. Move the transducer arrays 0.75 inches from the last areas where you placed them (so that the adhesive gel is between the red marks). If the reddening worsens: - Consult the study doctor |
| Blisters under the transducer arrays | Rare side effect | Ask the study doctor to prescribe you an antibacterial cream Use it as indicated by the doctor |
| Itching under the transducer arrays | Rare side effect | - When you change the transducer arrays, use an over-the-counter 0.1% hydrocortisone cream - Move the transducer arrays away from the last areas where you placed them (so that the adhesive gel is between the red marks). If the itching worsens: 1. Consult the study doctor |
| Pain under the transducer arrays | Rare side effect | Stop the treatment. Ask the study doctor. |

# ASSISTANCE AND INFORMATION

Technical support

- If you need technical support, contact the DSS. Contact information for your local DSS will be provided to you separately.

Clinical support

- If you notice any change in your health or any side effect from the treatment, call the study doctor at the site in charge of your treatment.

# TRAVELING WITH OPTUNE®

Optune® batteries contain material with lithium ions and cannot be checked as baggage on passenger aircraft. They can be carried in the passenger cabin.

When you travel to another country with the Optune device, use the proper electrical cable that has been supplied with the Optune® treatment kit.  Travel adapters cannot be used with the Optune® treatment kit.

If you have any questions about travel restrictions, contact your DSS.

Note: The Optune® device and INE Transducer Arrays will activate metal detectors.

# EMISSION OF RADIATION AND ELECTROMAGNETIC COMPATIBILITY

The Optune® treatment kit and the battery charger (ICH9100) and the power supply (SPS9100) that come with it require special electromagnetic compatibility (EMC) precautions and must be installed and operated according to the EMC information provided below.

Portable and mobile RF communications equipment may affect the Optune® treatment kit and the battery charger that comes with it.

The Optune® device must be used only with the following cables and accessories:

1. CAD9100 connection cable

2. INE9000 INE Transducer Arrays (sterile)

3. IBH9100 battery

4. SPS9100 power supply

5. ICH9100 charger

6. Unshielded AC power cords only for use indoors, with a maximum length below 3.0 m

The use of accessories, parts and cables other than those specified may increase EMISSIONS or reduce the IMMUNITY of the Optune® treatment kit.

Table 1: MANUFACTURER’S guide and declaration on ELECTROMAGNETIC EMISSIONS for all ME EQUIPMENT and ME SYSTEMS

| Manufacturer’s guide and declaration: electromagnetic emissions |  |  |
| --- | --- | --- |
| The Optune treatment kit is designed for use in the electromagnetic environment specified below. The customer or user of the Optune treatment kit must ensure that it is used in such environment. |  |  |
| Emissions test | Compliance | Electromagnetic conditions: guide |
| RF emissions CISPR 11 | Group 1 | The Optune treatment kit uses RF power only for its internal function. Therefore, its RF emissions are very low and the likelihood of it causing interference in nearby electronic equipment is very low. |
| RF emissions CISPR 11 | Class B | The Optune treatment kit is suitable for use in all establishments, including home use and in establishments directly connected to the low-voltage public power grid that supplies power to residential buildings. |
| Harmonic current emissions IEC 61000-3-2 | Class A |  |
| Voltage fluctuations and flicker emissions IEC 61000-3-3 | Compliant |  |

| Manufacturer’s guide and declaration: electromagnetic emissions |  |  |
| --- | --- | --- |
| The ICH9100 charger and the SPS9100 power supply have been designed for use in the electromagnetic environment specified below. The customer or user of the ICH9100 charger and the SPS9100 power supply must ensure that it is used in such an environment. |  |  |
| Emissions test | Compliance | Electromagnetic conditions: guide |
| RF emissions CISPR 11 | Group 1 | The ICH9100 charger and the SPS9100 power supply use RF power only for their internal function. Therefore, its RF emissions are very low and the likelihood of it causing interference in nearby electronic equipment is very low. |
| RF emissions CISPR 11 | Class B | The ICH9100 charger and the SPS9100 are suitable for use in all establishments, including home use and in establishments directly connected to the low-voltage public power grid that supplies power to residential buildings. |
| Harmonic current emissions IEC 61000-3-2 | Class A |  |
| Voltage fluctuations and flicker emissions IEC 61000-3-3 | Compliant |  |

Warning: The Optune® treatment kit, the ICH9100 charger and the SPS9100 power supply should not be used adjacent to or stacked with other equipment.

Table 2: MANUFACTURER’S guide and declaration on ELECTROMAGNETIC IMMUNITY for all ME EQUIPMENT and ME SYSTEMS

| Manufacturer’s guide and declaration: electromagnetic immunity |  |  |  |
| --- | --- | --- | --- |
| The Optune treatment kit, including the Optunedevice, the ICH9100 charger and the SPS9100 power supply, have been designed for use in the electromagnetic environment specified below. The customer or user of the Optune treatment kit must ensure that it is used in such environment. |  |  |  |
| Emissions test | IEC 60601 Testing level | Compliance level | Electromagnetic conditions: guide |
| Electrostatic discharge (ESD) IEC 61000-4-2 | ± 6 kV by contact ± 8 kV by air | ± 6 kV by contact ± 8 kV by air | Flooring must be hardwood, concrete or ceramic tiles. If the floor is covered with a synthetic material, the relative humidity should be at least 30%. |
| Electrical fast transient / burst transients IEC 61000-4-4 | ± 2 kV for power supply lines ± 1kV for input/output lines | ± 2 kV N/A |  |
| Overvoltage IEC 61000-4-5 | ± 1 kV line to line ± 2 kV line to ground | N/A |  |
| Voltage dips, short interruptions and voltage variations on power supply input lines IEC 61000-4-11 | UT <5% (>95% dip in UT) for 0.5 cycle 40% UT (60% dip in UT) for 5 cycles 70% UT (30% dip in UT) for 25 cycles UT <5% (> 95% dip in UT) for 5 s | N/A |  |
| Magnetic field of grid frequency (50/60 Hz) IEC 61000-4-8 | 3 A/m | 3 A/m | The magnetic fields of the grid frequency should be at levels characteristic of a typical location in a typical commercial or hospital setting. |
| NOTE: UT is the AC grid voltage before the test level is applied |  |  |  |

| Manufacturer’s guide and declaration: electromagnetic immunity |  |  |  |
| --- | --- | --- | --- |
| The ICH9100 charger and the SPS9100 power supply have been designed for use in the electromagnetic environment specified below. The customer or user of the ICH9100 charger and the SPS9100 power supply must ensure that it is used in such an environment. |  |  |  |
| Emissions test | IEC 60601 Testing level | Compliance level | Electromagnetic conditions: guide |
| Electrostatic discharge (ESD) IEC 61000-4-2 | ± 6 kV by contact ± 8 kV by air | ± 6 kV by contact ± 8 kV by air | Flooring must be hardwood, concrete or ceramic tiles. If the floor is covered with a synthetic material, the relative humidity should be at least 30%. |
| Electrical fast transient / burst transients IEC 61000-4-4 | ± 2 kV for power supply lines ± 1kV for input/output lines | ± 2 kV N/A | The quality of the power supply network should be that of typical commercial or hospital settings. |
| Overvoltage IEC 61000-4-5 | ± 1 kV line to line ± 2 kV line to ground | ± 1 kV line to line ± 2 kV line to ground | The quality of the power supply network should be that of typical commercial or hospital settings. |
| Voltage dips, short interruptions and voltage variations on power supply input lines Magnetic field of grid frequency (50/60 Hz) | UT <5% (>95% dip in UT) for 0.5 cycle 40% UT 3 A/m | UT <5% (>95% dip in UT) for 0.5 cycle 40% UT 3 A/m | The quality of the power supply network should be that of typical commercial or hospital settings. The magnetic fields of the grid frequency should be at levels characteristic of a typical location in a typical commercial or hospital setting. |
| NOTE: UT is the AC grid voltage before the application of the test level = 120 V and 230 V |  |  |  |

Normal operation: The Optune® treatment kit is working correctly when the blue LED surrounding the TTFields button is on and no notification signal sounds. The ICH9100 charger is working correctly when all the LEDs are on. The SPS9100 power supply is working properly when the blue LEDs surrounding the TTFields button in the Optune® treatment kit are on and no notification signal sounds.

Table 3: MANUFACTURER’S guide and declaration on ELECTROMAGNETIC IMMUNITY for all ME EQUIPMENT and ME SYSTEMS that are not LIFE-SUPPORTING

| Manufacturer’s guide and declaration: electromagnetic immunity |  |  |  |
| --- | --- | --- | --- |
| The Optune® treatment kit is designed for use in the electromagnetic environment specified below. The customer or user of the Optune treatment kit must ensure that it is used in such environment. |  |  |  |
| Emissions test | IEC 60601 test level | Compliance level | Electromagnetic conditions: guide |
| Conducted RF IEC 61000-4-6 Radiated RF IEC 61000-4-3 | 3 Vrms From 150 kHz to 80 MHz 3 V/m From 80 MHz to 2.5 GHz | 3 Vrms 3 V/m | Portable and mobile RF equipment should not be used at a shorter distance from any part of the Optune treatment kit, including the cables, than the recommended separation distance calculated using the equation applicable to the transmitter frequency. Recommended separation distance From 80 MHz to 800 MHz From 800 MHz to 2.5 GHz Where P is the maximum output power rating of the transmitter in watts (W), according to the manufacturer of the transmitter and d is the recommended separation distance in meters (m). Field strengths from fixed RF transmitters, as determined by an electromagnetic site surveya, should be less than the compliance levels in each frequency range. b) In the vicinity of equipment marked with the following symbol, interference may occur: |
| NOTE 1: At 80 MHz and 800 MHz, the higher frequency range applies. NOTE 2: These guidelines may not be applicable in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. |  |  |  |
| - Field strengths from fixed transmitters, such as base stations for radio (cellular/cordless) telephones and land mobile radios, amateur radio, AM and FM radio transmissions and TV transmissions cannot be predicted theoretically with accuracy. To assess the electromagnetic environment due to fixed RF transmitters, an electromagnetic site survey should be considered. If the measured field strength in the location where the Optune treatment kit is used exceeds the above applicable RF compliance level, the Optune treatment kit should be observed to ensure normal operation. If abnormal performance is observed, additional measures may be necessary, such as reorienting or repositioning the Optune treatment kit. - Above the frequency range of 150 kHz to 80 MHz, the field strengths should be less than 3 V/m |  |  |  |

Table 4: Recommended separation distances between portable and mobile RF communications equipment and the ME EQUIPMENT or ME SYSTEM, for ME EQUIPMENT and ME SYSTEMS that are not LIFE-SUPPORTING

| The Optune treatment kit is designed for use in an electromagnetic environment in which radiated RF disturbances are controlled. The customer or user of the Optune treatment kit can help prevent electromagnetic interference by keeping a minimum distance between the portable and mobile RF communications equipment (transmitters) and the Optune treatment kit as recommended below, in accordance with the maximum output power of the communications equipment. |  |  |  |
| --- | --- | --- | --- |
| Rated maximum output power of the transmitter in W | Separation distance according to transmitter frequency, in m |  |  |
|  | From 150 kHz to 80 MHz | From 80 MHz to 800 MHz | From 800 MHz to 5.5 GHz |
| 0.01 | 0.116 | 0.116 | 0.233 |
| 0.1 | 0.368 | 0.368 | 0.736 |
| 1 | 1.16 | 1.16 | 2.33 |
| 10 | 3.68 | 3.68 | 7.36 |
| 100 | 11.6 | 11.6 | 23.2 |
| In the case of transmitters that are not rated at a maximum output power listed above, the recommended separation distance d in meters (m) can be calculated through the equation applicable to the transmitter frequency, where P is the maximum output power rating of the transmitter in watts (W), according to the transmitter manufacturer. NOTE 1: At 80 MHz and 800 MHz, the separation distance of the higher frequency range applies. NOTE 2: These guidelines may not be applicable in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. |  |  |  |